## Performance Critera
| Performance Criteria | Acceptable Solutions | Met/Not Met | Proposed solution | PHRU comments
P1: | Functionality A private health facility must be suitably planned and fitted out to allow the effective and efficient functioning of the facility, taking into consideration- (a) the plan layout; and (b) adjacencies of suites, rooms and facilities; and (c) the suitability of room sizes where not nominated in the acceptable solution; and (d) administration and office facilities; and (e) accessibility to facilities by patients, staff, and others as required; and (f) patient privacy; and (g) location of furniture, fixtures, service outlets, and display panels. | A1: There are no corresponding acceptable solutions provided for the performance criteria P1. Queensland Health approval is required.
P2: | Emergency Power A private health facility must have access to emergency power to ensure the care and safety of patients, and the safety of staff at all times. | A2: A backup generator is provided in a private health facility to ensure uninterrupted power is supplied to- (a) intensive care units; and (b) operating theatres; and (c) coronary care units; and (d) procedure rooms; and (e) induction rooms; and (f) neonatal special care units; and (g) delivery suites; and (h) emergency departments; (i) at least 1 outlet for each clinical unit
P3: | Electrical Installations A private health facility must provide electrical components, fittings, appliances, equipment and apparatus to a standard that will provide a safe environment for patients, staff and visitors at all times. | A3: A private health facility provides electrical fittings, appliances, equipment and apparatus in accordance with AS 3003 and the relevant statutory Electricity Authority requirements.
P4: | Lighting A private health facility must provide lighting that will provide a safe environment for patients, staff and visitors at all times. | A4: A private health facility provides- (a) lighting in accordance with AS/NZS 1680.2.5; and
(b) night lighting in general wards and associated corridors in accordance with Section 10.1 of AS1680.2.5.
P5: | Medical Gases A private health facility must provide medical gases that will provide a safe environment for patients, staff and visitors at all times. | A5: A private health facility provides medical gases in accordance with AS 2896.
P6: | Storage of Records A private health facility must provide appropriate secure record storage to ensure patient confidentiality at all times. | A6: A private health facility stores all patient medical records and registers in a 1 hour (1010 deg. cent.) fireproof compartment in accordance with AS4390.6.
P7: | Air Conditioning A private health facility must provide appropriate air temperature control that will provide a safe and comfortable environment for patient, staff and visitors. | A7: A private health facility provides- (a) general comfort air conditioning in all clinical units; and
(b) air conditioning with specific requirements as follows - (i) separate air conditioning to clean areas (including sterile storage areas) with no common supply to systems serving other areas; and
(b) (ii) clean areas (including sterile storage areas) are positive air-pressured in relation to other areas so that the air flow pattern is out and away from these areas; and
(b) (iii) dirty areas are air conditioned and negative air-pressured so air flow from all other areas is into these areas; and
(b) (iv) all sterile areas including sterile stock storage areas have a two (2) stages air filtration system; and
(b) (v) operating theatres have HEPA filters (absolute with arrestance efficiency filters of 99.97%).
P8: | General Wards A general ward in a private hospital must facilitate the safety and care of patients, and the safety of staff, at all times taking into consideration - (a) model of care; and (b) floor area; and (c) access; and (d) facilities. | A8: A general ward in a private hospital - (a) is a clinical unit; and
(b) is located on a single level of the building; and
(c) provides access for patients to a sitting room, within close proximity to the ward, with a minimum floor area of 13m2; and
(d) provides a - (i) fixed lockable storage facility for medication and dangerous drugs; and
(d) (ii) clinical hand basin in each single or multiple bed room, or is located outside the room in close proximity to the room’s entry; and
(d) (iii) provides a hand basin in each ensuite, toilet or bathroom.
P9: | Patient Rooms A room for the purpose of patient accommodation in a private hospital must facilitate the safety and care of patients, and the safety of staff, taking into consideration- (a) floor area; and (b) access; and (c) facilities. | A9: In a private hospital- (a) a patient room has a minimum floor area of - (i) 15m2 for single bed accommodation; or
(ii) 22m2 for 2 bed accommodation; or
(iii) 33m2 for 3 bed accommodation; or
(iv) 11m2 per patient for accommodation for 4 of more maternity patients;
or (v) 10.75m2 per patient for accommodation for 4 or more non-maternity patients;
(b) a patient room- (i) has a door of 1.2m minimum width opening directly onto a corridor; and
(ii) does not have direct access to another patient’s room; and
(c) a patient room is provided with- (i) a minimum of 6 GPO’s to every bed location wall mounted nominally 1200mm above finished floor level; and
(ii) a wall mounted GPO elsewhere in the room 300 mm above finished floor level; and
(iii) a GPO and connection point for television; and
(iv) a connection point (GPO or other) for patients’ nightlight and reading light.
P10: | Nurse Calling System A private health facility must provide a call system that allows patients and staff to alert nurses and other health care staff in a discreet manner at all times. | A10: A private health facility provides a nurse call system that has - (a) visual and audible notification of general and emergency help needed by patients and staff; and
(b) a call point located beside each patient bed; and
(c) a call point located in each patient area such as lounge, holding / treatment / recovery bays, operating theatres, procedure rooms, labour rooms, nurseries, interview rooms and patient bathrooms, showers, toilets and change areas; and
(d) a call point for emergency use located in each clinical unit; and
(e) an annunciator panel/s, visible from all parts of the clinical unit, to display the point of origin of the call; and
(f) a display above the entry door to each room, visible from the corridor, to indicate the point of origin of the call; and
(g) 2 alarm sounds dedicated to nurse call and emergency call respectively, audible from all parts of the clinical unit, which activates simultaneously with the visible display, indicating the point of origin of the call; and
(h) no method of deactivation of the call other than at its source.
P11: | Staff Stations A clinical unit in a private health facility must provide a suitable area to facilitate the care and safety of patients. | A11: A clinical unit in a private health facility has a staff station located within it which is - (a) a minimum of 12m2 in floor area; and
(b) centrally located; and
(c) located to allow clear access to all patient areas and corridors at all times; and
(d) located to allow a clear view of a nurse call annunciator panel, in accordance with A10(e) at all times.
P12: | Clean and Dirty Utility Rooms A private health facility must provide suitable ancillary facilities for clean and dirty utilities for each clinical unit to ensure the safety and care of patients, and the safety of staff at all times. | A12: (a) A clean utility room is provided and dedicated to each clinical unit in a private health facility and has - (i) a minimum floor area of 10m2; and
(ii) a storage facility for sterile supplies, instruments, equipment and procedural trolleys and which may be used for the storage and recharging of pumps; and
(iii) coved splash-backs and sealed impervious bench-tops and cupboards; and
(iv) at least one utility sink; and
(v) a clinical hand basin.
(b) A dirty utility room is provided and dedicated to each clinical unit and has - (i) a minimum floor area of 10m2; and
(ii) vermin-proofed accesses and penetrations; and
(iii) a bed pan sanitiser, a bed pan and urinal rack, a utensil washer for bowls and dishes, and a rim-flushing sink or sluice; and
(iv) at least one utility sink; and
(v) a clinical hand basin; and
(vi) cupboards for the storage of disposable and cleaning items; and
(vii) coved splash-backs and sealed impervious bench-tops and cupboards; and
(viii) an exhaust fan capable of providing sufficient air changes to maintain a constant odour free and hygienic area; and
(ix) an access door fitted with a self-closing device that maintains the door in a closed position when not in use; and
(x) floors and wall finishes which are seamless, impervious or welded, and washable, the floors being non-slip and graded to fall to a floor waste.
P13: | Cleaning Facilities A clinical unit in a private health facility must have adequate facilities to enable the efficient and timely cleaning of the unit | P13: A separate lockable cleaners’ room is located within close proximity of a clinica+C64l unit and has - (a) a minimum floor area of 4m2; and
(b) a cleaner’s sink; and
(c) a storage facility for cleaning agents, supplies, equipment, disposal of flowers, vases, general waste and cleaning trolleys.
P14: | Storage In Clinical Units Clinical units in a private hospital must contain suitable storage facilities, which enable the efficient, and safe care of patients, and the safety of staff at all times. | A14: A private hospital has - (a) general storage facilities provided in all clinical units except general wards which are - (i) a minimum of 9m2 in floor area; and
(ii) located in a low traffic area within the unit with access to all patient areas; and
(b) equipment storage facilities provided in all general wards which are - (i) a minimum of 20m2 in floor area; and
(ii) located in a low traffic area within the ward with access to all patient areas (Refer to P31 for General Storage)
P15: | Corridors All types of corridors in a private hospital must be adequate to facilitate the care and safety of patients, and the safety of staff, taking into consideration - (a) dimensions; and (b) support aids. | A15: (a) A primary corridor in a private hospital has- (i) a minimum width of 2.2m clear of handrails and other obstacles; and
(ii) a handrail on one side.
(b) A secondary corridor in a private hospital has- (i) a minimum width of 1.8m clear of handrails and other obstacles; and
(ii) a handrail on one side.
(c) A tertiary corridor in a private hospital in accordance with AS 1428.1 and AS 1428.2 Subsection 6.5.
P16: | Sanitary Facilities Sanitary and bathing facilities in a private health facility must be adequate to facilitate the care and safety of patients, and the safety of staff, taking into consideration- (a) patient numbers; and (b) disabled use; and (c) location; and (d) staff; and (e) temperature control; and (f) special equipment; and (g) ease of use; and (h) finishes; and (i) safety of use; and (j) vermin and insect infestation; and (k) emergency access. | A16: Sanitary and bathing facilities in a private health facility are provided in accordance with the following: (a) General wards have - (i) ensuites which are of a minimum of 3.5m2 in floor area and capable of accommodating a patient and at least 1 person assisting; or
(ii) a minimum of 1 WC per 4 patients; and
(iii) a minimum of 1 shower, suitable for use with a mobile shower chair, per 6 patients.
(b) Each clinical unit has a minimum of 1 WC, shower and basin for use by patients with disabilities in accordance with AS 1428.1 and AS1428.3 where paediatric facilities are provided.
(c) Showers and baths are located to provide access to patients for whom they are intended.
(d) The following are provided for staff use: (i) 1 WC in each clinical unit.
(ii) 1 shower within the private health facility.
(e) Thermostatically controlled hot water installations are installed in accordance with AS/NZS 3666.1.
(f) A private hospital provides - (i) a minimum of 1 bathroom, capable of facilitating a mobile bath trolley or an assisted bath with a mechanical lifting device; and
(ii) seating for at least 1 person in the dressing areas of all shower cubicles and bathrooms.
(g) All bathrooms, ensuites, WCs and shower cubicles provide - (i) securely fitted grab rails and/or guide bars for the assistance of patients in accordance with AS1428.1 and AS1428.3 where paediatric facilities are provided; and
(ii) a flexible hand held shower unit for all showers other than those in mental health units.
(h) All bathrooms, ensuites, WCs, and shower cubicles have - (i) floors that are slip resistant and graded to fall to a floor waste; and
(ii) intersections of walls and floors are coved in continuous materials; and
(iii) floor and wall finishes which are seamless or welded, impervious, and washable.
(i) A shower is not installed over a plunge bath.
(j) All bathrooms, ensuites, WCs and shower cubicles are constructed to prevent the intrusion of insects or vermin.
(k) All bathrooms, ensuites, WCs and shower cubicles are accessible at all times (including in the event of an obstruction preventing normal door swing).
P17: | Sterilisation Facilities A private health facility must provide adequate facilities for cleaning, sterilisation and storage of equipment and instruments to ensure the care and safety of patients, and the safety of staff at all times. | A17: There are no corresponding acceptable solutions provided for the performance criteria P17. Queensland Health approval is required. (Refer to General Requirements (B))
P18: | Intensive Care Units An intensive care unit in a private hospital must adequately facilitate the care and safety of its patients, and the safety of staff, at all times taking into consideration- (a) floor area; and (b) staff station; and (c) nurse call system; and (d) sanitary and bathing facilities. | A18: An intensive care unit in a private hospital has - (a) a minimum floor area of - (i) 20m2 for single bed accommodation or isolation rooms; or
(ii) 18m2 per bed for multi-bed accommodation; and
(b) a centrally located staff station complying with A11; and
(c) a nurse call system complying with A10; and
(d) sanitary and bathing facilities complying with A16 (b), (c) and (d)(i); and
(e) access to the use of - (i) a clean and dirty utility room complying with A12; and
(ii) a cleaners’ room complying with A13; and
(f) a dedicated store, with a minimum floor area of 25m2, for storage of equipment and linen for use in the unit; and
(g) for each bed, a minimum of 8 Cardiac Protected Electrical Outlets and outlets for oxygen, air and suction; and
(h) a clinical hand basin accessible from each bed.
P19: | Coronary Care Units A coronary care unit in a private hospital must adequately facilitate the care and safety of patients, and the safety of staff at all times, taking into consideration- (a) floor area; and (b) staff station; and (c) nurse call system; and (d) sanitary facilities; and (e) ancillary services; and (f) storage; and (g) special equipment; and (h) staff facilities. | A19: A coronary care unit in a private hospital has- (a) a minimum floor area of 18m2 per bed; and
(b) a centrally located staff station complying with A11; and
(c) a nurse call system complying with A10; and
(d) sanitary and bathing facilities complying with A16 (b), (c) and (d)(i); and
(e) access to the use of- (i) a clean and dirty utility room complying with A12; and
(ii) a cleaners’ room complying with A13; and
(f) a dedicated store, with a minimum floor area of 25m², for storage of equipment and linen for use in the unit; and
(g) for each bed, a minimum of 8 Cardiac Protected Electrical Outlets and outlets for oxygen, air and suction; and
(h) a clinical hand basin accessible from each bed.
P20: | Emergency Departments An emergency department in a private hospital must adequately facilitate the care and safety of patients, and the safety of staff, at all times, taking into consideration- (a) nurse call system; and (b) ancillary services; and (c) storage; and (d) staff change facilities; and (e) sanitary facilities; and (f) design layout; and (g) spread of infection. | A20: An emergency department in a private hospital includes - (a) a nurse call system complying with A10; and
(b) access to the use of - (i) a clean and dirty utility room complying with A12; and
(ii) a cleaners’ room complying with A13; and
(c) access to storage room for the storage of equipment and hardware used in the department; and
(d) a staff change room in accordance with the Workplace Health and Safety Act & Regulation 1995; and
(e) a minimum of 1 WC and 1 hand-basin for each sex, and 1 unisex disabled WC provided for patient use, in close proximity to the department. There are no corresponding acceptable solutions provided for the performance criteria P20(f) and P20(g). Queensland Health approval is required. (Refer to General Requirements (C))
P21: | Inductions Rooms A private hospital, which provides for the reception, care and treatment of maternity cases must adequately facilitate the care and safety of patients, and the safety of staff, at all times taking into consideration - (a) accommodation; and (b) special equipment; and (c) staff facilities; and (d) finishes. | A21: An induction room, where provided in a private hospital has - (a) a minimum floor area of 15m2 including a scrubbing area; and
(b) for each bed, a minimum of 8 Cardiac Protected Electrical Outlets and outlets for oxygen, air and suction; and
(c) a clinical hand basin; and
(d) floor and wall finishes complying with A16 (h).
P22: | Delivery Suites A delivery suite in a private hospital must adequately facilitate the care and safety of obstetric patients, and the safety of staff, at all times taking into consideration- (a) accommodation; and (b) acoustic privacy; and (c) special equipment; and (d) staff facilities; and (e) ancillary services; and (f) storage; and (g) finishes; and (h) staff station; and (i) nurse call system; and (j) special care facilities. | A22: A delivery suite in a private hospital has - (a) at least 1 delivery room that has - (i) a minimum floor area of 27m2 exclusive of cabinets; and
(ii) a scrub-up facility adjacent to the entry to each delivery room; and
(iii) an ensuite of sufficient size to allow for fully assisted use of all facilities within close proximity to the suite; and
(b) an STC rating of 50 for walls floors and ceilings; and
(c) a minimum of 8 Cardiac Protected Electrical Outlets, 2 oxygen outlets, 1 medical air outlet, 1 nitrous oxide outlet and 2 suction outlets; and
(d) a clinical hand basin; and
(e) access to the use of: (i) a clean and dirty utility room complying with A12; and
(ii) a cleaners’ room complying with A13; and
(f) a dedicated store, with a minimum floor area of 11 m², for the storage of equipment and linen used in the suite; and
(g) floor and wall finishes complying with A 16 (h); and
(h) a staff station complying with A11; and
(i) a nurse call system complying with A10; and
(j) an infant resuscitation area, with a minimum floor area of 3.72m², within the room or immediately adjacent to it.
P23: | Mature Infant Nurseries A nursery in a private hospital must adequately facilitate the care and safety of neonatal patients, and the safety of staff, at all times taking into consideration- (a) accommodation; and (b) staff facilities; and (c) visual access; and (d) special equipment; and (e) emergency call; and (f) special facilities; and (g) ancillary services; and (h) storage facilities. | A23: A mature infant nursery in a private hospital has - (a) a minimum floor area of - (i) 15m2; or
(ii) for units greater than 12 obstetric beds - 1.5m2 per bassinet and 0.9m clear and unobstructed passageway between each bassinet. 2 bassinets to every 4 obstetric beds; and
(b) at least 1 clinical hand basin; and
(c) clear glazed partitions installed complying with AS 1288; and
(d) 1 suction and 1 oxygen outlet and a minimum of 4 GPO’s per 4 bassinets (complying with AS 3003); and
(e) an emergency call system complying with A10(d); and
(f) a dedicated area within or adjacent to the nursery to allow easy examination and changing of the baby, and storage of necessary linen and equipment; and
(g) a dirty utility room complying with A12(b); and
(h) a storage facility complying with A14(a) which also may be used for the preparation of baby formulae.
P24: | Neonatal Special Care Units A neonatal special care unit in a private hospital providing for the reception, care and treatment of obstetric patients must adequately facilitate the care and safety of patients, and the safety of staff, at all times, taking into consideration - (a) accommodation; and (b) special equipment; and (c) emergency call; and (d) staff facilities; and (e) visual access; and (f) acoustic privacy; and (g) storage facilities; and (h) special facilities. | A24: A neonatal special care unit in a private hospital is a separate unit to the mature infant nursery and has - (a) an area allocated to incubator space at the rate of 5m2 per incubator exclusive of any area occupied by benches, furniture or other apparatus (fixed or mobile); and
(b) a minimum of 8 Cardiac Protected Electrical Outlets per incubator and a minimum of 3 oxygen, 3 suction & 3 medical air outlets and an individual examination light for each incubator; and
(c) an emergency call point complying with A10(d); and
(d) a scrub-up sink located in a dedicated area at the entry to the unit and a clinical hand basin accessible from each cot; and
(e) clear glazed partitions installed to comply with AS 1288 to allow a clear and unobstructed view for patients and staff; and
(f) an STC rating of 50 for walls, floors and ceilings; and
(g) a dedicated store for equipment and linen for use in the unit; and
(h) a dedicated area for staff to monitor the care of the patient; and
(i) a dirty utility area complying with A12(b).
P25: | Neonatal Intensive Care An intensive care nursery unit in a private hospital providing for the reception, care and treatment of obstetric patients must adequately facilitate the care and safety of patients, and the safety of staff, at all times, taking into consideration - (a) accommodation; and (b) special equipment; and (c) emergency call; and (d) staff facilities; and (e) visual access; and (f) acoustic privacy; and (g) storage facilities; and (h) special facilities; and (i) ancillary services. | A25: An intensive care nursery unit in a private hospital has - (a) an area allocated to incubator space at the rate of 8m² per incubator exclusive of any area occupied by benches, furniture or other apparatus (fixed or mobile); and
(b) a minimum of 8 Cardiac Protected Electrical Outlets per incubator and a minimum of 3 oxygen, 3 suction & 3 medical air outlets and an individual examination light for each incubator; and
(c) an emergency call point complying with A10(d); and
(d) a scrub-up sink located in a dedicated area at the entry to the unit and a clinical hand basin accessible from each cot; and
(e) clear glazed partitions installed to comply with AS 1288 to allow a clear and unobstructed view for patients and staff; and
(f) a STC rating of 50 for walls, floors and ceilings; and
(g) a dedicated store for equipment and linen for use in the unit; and
(h) a dedicated area for staff to monitor the care of the patient; and
(i) a dirty utility area complying with A12(b).
P26: | Mental Health Unit A mental health unit in a private health facility must provide for the reception, care and treatment of mental health patients and must adequately facilitate the care and safety of patients, and the safety of staff, at all times. | A26: A mental health unit in a private health facility is: (a) A clinical unit.
(b) There are no corresponding acceptable solutions provided for the performance criteria P26. Queensland Health approval is required. (Refer to General Requirements (G))
P27: | Operating Suites A private health facility in which surgical operations are performed must adequately facilitate the care and safety of patients, and the safety of staff, at all times, taking into consideration- (a) size; and (b) special equipment; and (c) finishes; and (d) emergency call; and (e) staff facilities; and (f) storage; and (g) ancillary services; and (h) sterile environment; and (i) administration; and (j) office facilities. | A27: A private health facility has an operating suite which has - (a) an operating theatre with a minimum floor area of - (i) 36m2 for general surgery; and
(ii) 42m2 for orthopaedic surgery; or
(iii) 50m2 for cardio-vascular surgery; and
(b) each operating theatre has - (i) anaesthetic gas scavenging system; and
(ii) 2 oxygen outlets, a nitrous oxide outlet, a medical air outlet and 2 suction outlets; and
(iii) a minimum of 12 Cardiac Protected Electrical Outlets; and
(iv) a scrub-up facility of a minimum floor area of 5m2 per theatre immediately adjacent to the theatre; and
(c) floor and wall finishes that are seamless, impervious, welded, smooth, non slip and graded where necessary to fall to floor waste. Intersection joins are coved in continuous materials; and
(d) a nurse call system complying with A10; and
(e) a staff change room and rest and recreation facilities in accordance with Workplace Health and Safety Act & Regulation 1995; and
(f) a storage room, associated with the theatre, with a minimum floor area of 25m², for the storage of linen, equipment and hardware used in the theatre; and
(g) includes access to the use of - (i) a clean and dirty utility room complying with A12; and
(ii) a cleaners’ room complying with A13. There are no corresponding acceptable solutions provided for the performance criteria P27(h), (i) and (j). Queensland Health approval is required. (Refer to General Requirements (B))
P28: | Procedure Areas An area in a private health facility in which specified clinical procedures are performed must facilitate the care and safety of patients, and the safety of staff, with regard to- (a) size; and (b) special equipment; and (c) finishes; and (d) emergency call; and (e) staff facilities; and (f) storage; and (g) ancillary services; and (h) decontamination area; and (i) sterile environment; and (j) administration; and (k) office facilities. | A28: A procedure area for the performance of specified clinical procedures is provided in a private health facility and has - (a) minimum floor areas of – (i) for endoscopy procedures: 30m2 exclusive of any fixtures; or
(ii) for cardiac catheter procedures: 45m2 exclusive of any fixtures; or
(iii) for procedures performed under more than simple sedation: 30m2 exclusive of any fixtures; or
(iv) for procedures performed under simple sedation: 25m2 exclusive of any fixtures; or
(v) for renal dialysis treatments: 8m2 per chair / trolley; and
(vi) for chemotherapy treatments: 8m2 per chair / trolley; and
(vii) for cardiac stress testing: 14m2 exclusive of any fixtures; and
(b) equipment and facilities as follows - (i) an anaesthetic gas scavenging system where more than simple sedation is performed; and
(ii) a minimum of 2 oxygen outlets and 2 suction outlets; and
(iii) 6 Body Protected Electrical Outlets except for cardiac catheter procedure room that requires 6 Cardiac Protected Electrical Outlets; and
(iv) a clinical hand basin; and
(v) a scrub-up facility as for A27(b)(iv) adjacent to the cardiac catheter procedure room and a procedure room where more than simple sedation is performed; and
(c) floor and wall finishes complying with A27(c); and
(d) an emergency call system complying with A10; and
(e) access to an area complying with A27(e); and
(f) a storage room for the storage of equipment and hardware used in the procedure room; and
(g) access to the use of - (i) a clean and dirty utility room complying with A12; and
(ii) a cleaners’ room complying with A13; and
(h) A dedicated room with direct access from the procedure room, for the processing, cleaning and disinfection of endoscopic instruments complying with AS 4187. There are no corresponding acceptable solution provided for the performance criteria P28(i), (j) and (k). Queensland Health approval is required. Refer to General Requirements (B).
P29: | Recovery A private health facility in which surgical operations and clinical procedures are performed under more than simple sedation must adequately facilitate the care and safety of patients recovering from anaesthesia, and the safety of staff, at all times, taking into consideration- (a) accommodation (b) special equipment; and (c) nurse call; and (d) staff facilities; and (e) storage facilities; and (f) ancillary services; and (g) secondary recovery. | A29: (a) There are no corresponding acceptable solution provided for the performance criteria P29(a). Queensland Health approval is required. Refer to General Requirements (D). A recovery unit in a private health facility has- (b) a minimum of 2 Body Protected Electrical Outlets and a minimum of 1 oxygen and 1 suction outlet; and
(c) a nurse call system complying with A9; and
(d) a minimum of at least 1 clinical hand basin for every 4 patient recovery cubicles; and
(e) access to a storage room for the storage of equipment and hardware used in the department; and
(f) access to a clean and dirty utility room complying with A12.
(g) If a private health facility provides a secondary recovery area, this area has - a. direct access to a patient change room; and
b. direct access to the disabled toilets and shower facilities; and
c. sanitary and bathing facilities complying with A16.
P30: | Day Hospital - General A day hospital must adequately facilitate the care and safety of patients, and the safety of staff, at all times, taking into consideration - (a) administration facilities; and (b) interviewing / examination room; and (c) staff facilities; and (d) visitors’ facilities; and (e) emergency exit; and (f) ancillary services; and (g) special equipment; and (h) storage; and (i) nurse call system; and (j) food preparation. | A30: A day hospital has - (a) a dedicated area complete with reception and clerical facilities for separate charting and recording of admission details of patients by administrative staff and accommodation for storage of files complying with A6; and
(b) a dedicated room for the interviewing and/or examination of patients in privacy having - (i) curtains and/or privacy screens; and
(ii) a clinical hand basin; and
(c) staff facilities including - (i) a staff station in accordance with A11; and
(ii) a staff change room and rest and recreation facilities in accordance with the Workplace Health and Safety Act & Regulation 1995; and
(iii) a clinical hand basin located in all patient care areas; and
(d) a minimum of 1 WC and hand-basin for each sex and one unisex disabled toilet, in close proximity for use by visitors; and
(e) an emergency exit allowing access from any operating room, recovery area or other area; and
(f) access to the use of - (i) a clean and dirty utility room complying with A12; and
(ii) a cleaners’ room complying with A13; and;
(g) suction and oxygen storage located to allow quick and easy access by staff for patient use in emergency and non emergency situations; and
(h) storage facilities complying with A32 and a separate rooms for the storage of flammable goods; and
(i) a nurse call system complying with A10; and
(j) a food preparation facility complying with A33.
P31: | Day Hospitals - Surgical And Endoscopy Procedures A day hospital at which surgical and endoscopy procedures are performed must adequately facilitate the care and safety of patients, and the safety of staff, at all times. | A31: A day hospital at which surgical and endoscopy procedures are performed, has - (a) facilities complying with A30; and
(b) an operating theatre complying with A27 and a procedure room complying with A28; and
(c) a patient change room equipped with lockers and providing direct access to disabled toilets and shower facilities; and
(d) a pre-operative waiting area with direct, access to the operating theatre and procedure room; and
(e) a recovery area complying with A29; and
(f) sanitary and bathing facilities complying with A16
P32: | Storage - General A private health facility must have adequate storage facilities to ensure the care and safety of patients, and the safety of staff, at all times. | A32: Where not otherwise provided for in this Code, a private health facility has - (a) a general storage room, for the storage of bulk supplies from the Stores Unit and Sterile Supply Unit, which - (i) has a minimum floor area of 9m2; and
(ii) is fitted with a dead lock and located in a low traffic area; and
(b) a disposal room, for the storage of waste and used linen, which - (i) has a minimum floor area of 9m2; and
(ii) is located in close proximity to all functional areas; and
(iii) is accessible from within the unit and externally; and
(iv) is fitted with a dead lock. (Refer to 14 for Storage in Clinical units)
P33: | Kitchens A private health facility must have food preparation and food storage facilities adequate to ensure the care and safety of its patients, and the safety of staff and visitors, at all times, taking into consideration - (a) kitchen area; or (b) servery area. | A33: A private health facility provides - (a) a kitchen area for catering purposes which complies with the Food Hygiene Regulation 1989; or
(b) a servery area, that is not in a kitchen, which complies with the Food Hygiene Regulation 1989 and has - (i) a workbench with an impervious top and splashback; and
(ii) a single or double bowl stainless steel sink set in the benchtop supplied with hot and cold reticulated water, lever action (with off at wall) or automatically activated taps, a soap dispenser, a towel dispenser and a glove dispenser; and
(iii) a hand basin, if a single bowl sink is fitted under (ii).
(c) Each clinical unit must have access to a kitchen or servery.
P34: | Laundries A private health facility must provide laundry services that will adequately ensure the care and safety of patients, and the safety of staff, at all times. | A34: A private health facility - (a) has a laundry that is - (i) located so that any entry or egress is not through a room used in the preparation or storage of food stuffs; and
(ii) in accordance with AS 4146; or
(b) if a laundry is not provided, has a dedicated room for the storage of dirty linen which has- (i) ventilation in accordance with A12.(b)(viii); and
(ii) a clinical hand basin

## Legend
Met
Met with endorsed variation
Not Met
Advice required
N/A
Clarify